## Slide 1: Instrument API and XML Method Modification Interface A brief introduction
Tony Zhao and Jesse Canterbury

## Slide 2: Overview
Instrument Application Programming Interface (IAPI)
  Introduction
  Architecture
  Example
XML Method Modification Interface (XMMI)
  Architecture
  Example
Notes: Overview of the first half of today
Some special guest presenters on their experience using these tools

## Slide 3: Overview
Instrument Application Programming Interface (IAPI)
  Introduction
  Architecture
  Example
XML Method Modification Interface (XMMI)
  Architecture
  Example
Notes: Overview of the first half of today
Some special guest presenters on their experience using these tools

## Slide 4: Why an API?
“I wish I could do …”
“I have this awesome new experiment …”
“Can I do this in method editor?”
“I only want to select peaks that have odd masses …”
Notes: Why an API?
Lots of questions about doing cool things without mass spectrometers

## Slide 5: Why an API?
“I have this awesome new experiment …”
“I only want to select peaks that have odd masses …”
Users can only control the Mass Spectrometer through our Tune application and Method acquisition.
Although both provide great flexibility in acquisition strategies they don’t encompass all possibilities.
Providing an Application Programming Interface enables users to extend the capabilities of Thermo Scientific™ Orbitrap™ mass spectrometers.
Allows real-time data acquisition and control.
Can be used with Thermo Scientific™ Exactive™ and Thermo Scientific™ Tribrid™ MS systems
“I wish I could do …”
“Can I do this in method editor?”
Notes: We provide both Tune and Method Editor to allow flexibility, but not everything is possible
By providing an API, we can extend the capabilities of the MS
It also allows real-time control of the MS

## Slide 6: Instrument API Git Repository
  https://github.com/thermofisherlsms/iapi
Public interfaces for
  Thermo Scientific Tribrid MS
  Thermo Scientific Exactive MS
Documentation
  Previous ASMS Posters
  Help files
VS Solution of Example Applications
External Requirements
Tune Installation
.NET 4.6.2+
License Agreement in place
IAPI License Key Activated
Notes: Again, the IAPI is hosted on our GitHub page
It contains all the interfaces and documentation to get started
You’ll need a few things to use it

## Slide 7: Interfaces
Thermo Scientific Exactive Series
Thermo Scientific Tribrid Series
IAPI Version
Instrument Extensions
Misc.
All interfaces are in the git repository, in the lib/ directory
Notes: We have two IAPI-enabled instruments lines
The Exactive series came first with the first API definitions
  It is broken into multiple assemblies.
The Tribrid series came next, with a newer API definition
  Here again, we have the extensions for the Fusion API

## Slide 8: Differences and Similarities between Exactive and Tribrid Series IAPI
Although IAPI 1.X (supports Exactives) is not binary compatible with IAPI 2.X (supports Tribrids), they share considerable amount of structure
  Namespaces and hierarchies are virtually unchanged
  A few name changes or minor reorganization
  Removal of some interfaces
  Spectrum-related interfaces moved into own assembly
The logical flow for both receiving scan data and sending scans definitions to the instrument is identical
Different methods for instantiating the IAPIs
  Tribrid-series uses a Factory method model with Microsoft’s Managed Extensibility Framework (MEF)
  Exactive-series uses the System.Reflection.Assembly API
Notes: I would like to highlight some of the differences between the exactive and tribrid APIs

## Slide 9
  The rest of this presentation will focus on the Tribrid-series IAPI.
  Most topics are applicable to the Exactive-series with minor differences.

## Slide 10: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started
Notes: A lot of you probably don’t know the architecture of our mass spec system. So I will briefly walk through different components.

## Slide 11: MS Architecture
Mass Spectrometer
Data System
Local Switch
Tune Application
Instrument
Service
Rawfile
Notes: Here is the basic lay out of the MS architecture
The ms talk to data system over network/a switch. The data system runs the instrument service. The instrument service provide data to both the tune app, but also write to the raw file

## Slide 12: IAPI Architecture
Mass Spectrometer
Data System
Local Switch
Tune Application
Instrument
Service
3rd Party
Application
IAPI-Enabled App
Rawfile
Notes: So the API is basicly integrated into this instrument service.
The 3rd party app you guys write can talk to the API. And downstream everything is the same

## Slide 13: IAPI Data and Control Flow
Handles scan Event
Processes the data
Do whatever you want!
Displays scan in Tune
Writes to Rawfile
Forwards data to API
Mass Spectrometer
Data System
Instrument
Service
3rd Party
Application
Scan Data
Stream
Converts data into IAPI format
Raises Event to any subscribers
Notes: This is kind of a summary of the previous slice.
In terms of control flow, it has 3 layers, the instrument service, the API and the your application.
The mass spec is generating some spectra, it first reach instrument service which displays the scan in Tune, it writes to RawFile and Forwards data to API.
Now inside API, it kind of converts data into IAPI format, and it will raise events to any subscribers that listen on the IAPI side.
The application layer will have an event handler that handles the scan event data, and basicaly do whatever you want.

## Slide 14: IAPI Data and Control Flow
Calls IAPI to create IScanDefinition object
Changes definition as sees fit
Handles scan Event
Processes the data
Do whatever you want!
Displays scan in Tune
Writes to Rawfile
Forwards data to API
Mass Spectrometer
Data System
Instrument
Service
3rd Party
Application
Scan Data
Stream
Converts data into IAPI format
Raises Event to any subscribers
Validates scan definition
Forwards scan to Instrument Service
Scan Definition
Logs event
Downloads scan definition to instrument
Notes: That’s the basic data flow, now if you want to change what the mass spec is doing, maybe after the event, you can generate a IScanDefinition object in the IAPI, and you modify that definition whatever you sees fit.
You send it thru the API, that validates the scan definition, and then forwards it into the instrument service.
Instrument service logs the scan event, and downloads it to the instrument. Here I show a cut-out image of the scanDefinition table in Tune, to kind of show what a scan definition looks like.
And that will influence how instrument is taking scans.

## Slide 15: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started
Notes: Lets dig in deeper into the receiving side of things

## Slide 16: MS Scan Interfaces
IMsScanContainer exposes IMsScans
IMsScan contains metadata and derives from ISpectrum
ISpectrum contains a collection of ICentroids
ICentroids contains meta data and derives from IMassIntensity
IMassIntensity is the m/z and intensity data
Notes: This show layers of interfaces in C#. I stands for interface. It shows the hiarachy from the user level down to the low level.
Ultimately you have a MSScan container, that has a event called MsScanArrived, that’s what you should subscribe to. You can have multiple listeners to that scan event.
When you get that event, you will get a copy of MSScans. A MsScan represents both spectrum and all the meta data with that spectrum, so like the Trailer data, the Header data the TuneData and anything like that.
So MSScan is derived from ISpectrum. It will contain a list of Centroids among other things. Centroids consist of a bunch of Icentroid, which contain charge, resolution, is monoisotopic. And Icentroid is derived from IMAssIntensity.

## Slide 17: Scan Data Stream
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Receiving the scan data is through the Event Driven programming model
Scans are sent via the .NET Event and EventHandler mechanism
  Multiple handlers per scan can be registered
  Minimal overhead and latency
Receives Scans from both Tune and Method acquisitions
Notes: This example is in GitHub, If you want to modify and look at it.
This show the bare minimum of the code behind to write sth that subscribe to MS scan event. This is an event driven programming model.
IFusionMsScanContainer is the basic API to get MsScan, there is only 1 of them.
Then here is a console app that basicly just waits util the key is escape.
Then the key is S means Subscribe to the MSScanEvent. So I register this call-back function. Or the key is U I unsubscribe that. I will show a demo in the next slice. And if it is other key, I default here.
What subscribe does is, this MsScanEvent, which has a function called getScan. I just print out from header the ScanNumber.

## Slide 18: Scan Data Stream
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Receiving the scan data is through the Event Driven programming model
Scans are sent via the .NET Event and EventHandler mechanism
  Multiple handlers per scan can be registered
  Minimal overhead and latency
Receives Scans from both Tune and Method acquisitions
tSIM running ~21 Hz (47.5 ms) from Tune.
  Average time of 48.4 ms between scans received (n = 200)
Notes: Here I show you how you are running it. Now I hit S, u can see how fast it is, there is really no latency. Now I hit U to unsubscribe. So it doesn’t threat what actually is running, it just kind of listening to the data.

## Slide 19: Scan Data Stream
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Receiving the scan data is through the Event Driven programming model
Scans are sent via the .NET Event and EventHandler mechanism
  Multiple handlers per scan can be registered
  Minimal overhead and latency
Receives Scans from both Tune and Method acquisitions
tSIM running ~21 Hz (47.5 ms) from Tune.
  Average time of 48.4 ms between scans received (n = 200)
  The IAPI is responsive and has no impact on the MS acquisition
Notes: This is how I justify the latency. I ran a tSIM that was running at on average 47.5ms per scan. On the receiving side (Console app), it takes at 48.4ms between scans. So 1ms on average of latency

## Slide 20: Scan Data Contents
Notes: So I took that MSScanArrived handler, I put a break point here when I was running the program.

## Slide 21: Scan Data Contents
Notes: This just shows the data that contains there. All the mz peaks.
I just want to show, hey we gave you a lot of data, and you can do a lot with it. And it’s relatively easy to explore.

## Slide 22: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started

## Slide 23: IScans Interface
Sending scans to the instrument is through the IScans interface
  Provides methods for creating different types of IScanDefinitions
  Provides methods for setting and cancelling scans.

## Slide 24: IScanDefinition
Scans are defined in the IScanDefinition interface
  The RunningNumber property is for general purpose use
    It shows up in the Scan Trailer as the ‘Access ID’ field
  The Values property is a Dictionary<string,string>
    Keys are the scan parameters
    Values are their associated values
    Values can take different formats based on the parameter
    Invalid, or nonsense values, are ignored
Example of MS3 of 524@35 191@25
Scan 150 – 550 m/z

## Slide 25: IScanDefinition

## Slide 26: Scan Management
Mass Spectrometer
Data System
Local Switch
Instrument
Service
3rd Party
Application
IAPI-Enabled App
Sending Scan to MS

## Slide 27: Scan Management
Mass Spectrometer
Data System
Local Switch
Instrument
Service
3rd Party
Application
IAPI-Enabled App
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Sending Scan to MS

## Slide 28: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 29: Scan Management - ICustomScan
Single scan, drops to lower level after completion
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 30: Scan Management - ICustomScan
Single scan, drops to lower level after completion
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Once executed, waits the Single Processing Delay (s) for a new Custom scan to arrive*
Default delay is 0 seconds (no waiting)
Immediately ends the wait if a new Custom scan arrives or there are pending ones
  The Tribrid uses a circular buffer (40-capacity) to store them
* Single Processing Delay is not enabled when occurring during Tribrid method acquisition, due to acquisition pipeline

## Slide 31: Scan Management - IRepeatingScan
Repeats scan until explicitly canceled by user
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Continually runs the same scan until:
  Cancelled by the user
  Replaced by a new Repeating Scan
  Preempted by a Custom Scan

## Slide 32: Scan Management - IRepeatingScan
Runs Scans and logic defined in the method
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Runs the normal method logic until completion or new IAPI scan arrives

## Slide 33: Scan Management - IRepeatingScan
Runs the Scan defined in Tune
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 34: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method

## Slide 35: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating

## Slide 36: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom

## Slide 37: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating

## Slide 38: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating

## Slide 39: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating
Method Completed

## Slide 40: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating
Method Completed

## Slide 41: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 42: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Method Scans
IAPI Scans
Method Scans

## Slide 43: Starting Acquisitions from the IAPI
The IAcquisition interface enables the IAPI to start/stop acquisitions
It also has Events for listening to when an acquisition starts and stops

## Slide 44: Starting Acquisitions from the IAPI
The IAcquisition interface enables the IAPI to start/stop acquisitions
It also has Events for listening to when an acquisition starts and stops

## Slide 45: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started

## Slide 46: License Activation
Access to the IAPI is given after a fully executed IAPI License is in force.
Contact info.IAPI@thermofisher.com to get started
  https://github.com/thermofisherlsms
  Please include which IAPI (Tribrid or Exactive) you would like access to in your initial correspondence
Notes: To get start with it, we have this ThermoMSLicesing account. Or the easiest way would be to go to our github page. There needs some paper work

## Slide 47: License Activation
Access to the IAPI is given after a fully executed IAPI License is in force.
Contact info.IAPI@thermofisher.com to get started
  Please include which IAPI (Tribrid or Exactive) you would like access to in your initial correspondence
A license key will be provided to you and is entered in Tune’s “About Tune” dialog
Restart both the Instrument and Data system computer and the API will be enabled
Activating the License on the Tribrids
Notes: After signing the license, we will provide with a licence key according to your instrument ID number. And you can just add license here.

## Slide 48: Overview
Instrument Application Programming Interface (IAPI)
  Introduction
  Architecture
  Example
XML Method Modification Interface (XMMI)
  Architecture
  Example
Notes: Overview of the first half of today
Some special guest presenters on their experience using these tools

## Slide 49
Method courtesy of Christopher Rose
The Method Editor offers tremendous flexibility in method creation;
sometimes manually editing complex methods is laborious.

## Slide 50: Programmatically Modify Fusion Methods
Method File
Modifications
Modified
Method File
XML Method
Modification Interface

## Slide 51: Method Modifications Git Repository
  https://github.com/thermofisherlsms/meth-modifications
Supports Fusion and Triple-Quad Methods
  Versions 1.2, 2.0, 2.1, 3.0, 3.1, 3.3
Documentation
Complete Examples
VS Solution
  Command Line Program
  Heavily commented code
Schemas
External Requirements
Tune Installation
  Full or Workstation
.NET 4.6.2+

## Slide 52: XML Modification Interfaces
Notes: Here I show the two interfaces we use in XML Modification.
Thermo.TNG.MethodXMLFactory.dll
Thermo.TNG.MethodXMLInterface.dll
To compile, you need to include both assemblies.

## Slide 53: Complete Code Example
Notes: Top two line are including assemblies.
The following line 4 to 6 are templateMethod, modificationXML, and the outputMethod.
Following we specify the Model and version.
Here we create the context given the Model and Version. We open the template, apply the modification, and eventually save it.

## Slide 54: Method Modifications Visual Studio Solution
The XmlMethodChanger Solution contains two projects
  XmlMethodChanger.lib
    A public, static class library (MethodChanger)
    Wraps the two Thermo.TNG.*.dll assemblies
    Provides public methods for common tasks
    Serves as an example of the API Usage
  XmlMethodChanger.cmd
    Command-line program that uses XmlMethodChanger.lib
    Uses third-party NuGet package (Command Line Parser Library)
    Can be used as a standalone tool for modifying methods
Notes: This is a solution in GitHub. A command line program for running that. It consists of two projects, one is the library. It kind of just wrap the two assemblies. Project 2 is the .cmd, the command line prgoram that calls the library to do the work, provide an interface to the command line.

## Slide 55: Demonstration
The “template” method contains the structure of the experiment
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository
Notes: This is a sample in Github down here. This is a template method, I just made a targeted mass exclusion list that has mothing in there.

## Slide 56: Demonstration
The “template” method contains the structure of the experiment
The modification XML specifies a list of transformations to perform on the “template” method file
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository
Notes: Here I write the XML modification.
I’m adding a couple of entries, 195, 262 to this targeted mass exclusion.
For the targetedMassInclusion, I’m adding the same mass but at a different intecity threshold, 5e7 and 1e2.

## Slide 57: Demonstration
The “template” method contains the structure of the experiment
The modification XML specifies a list of transformations to perform on the “template” method file
The transformed method can then be saved to a separate .meth file
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository

## Slide 58: Demonstration
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository
Notes: I load the template in, there is nothing in there. Now I ran the bat script, which will call the command line to modify the template based on the XML. It generates the new one, I load it into the methodEditor, and you can see it loads the new data according to our numbers in xml.

## Slide 59: THANK YOU!