## Slide 1: Instrument API and XML Method Modification Interface
ASMS Orbitrap Power User Meeting – June 2nd, 2018, San Diego, CA
Derek Bailey, Ph.D.
Notes: Thank Shannon for spearheading
Thank attendees for coming

## Slide 2: ASMS 2018 Warm Up Slide
https://xkcd.com/1168/
Notes: Many slides over the next few days at ASMS so I’ll start with a warm up
With an obligatory XKCD comic
Once someone programs AI to search stackoverflow.com for answers, I’m out of a job…

## Slide 3: Overview
Instrument Application Programming Interface (IAPI)
  Real-time control and acquisition of Orbitrap Mass Spectrometers using .NET
XML Method Modification Interface (XMMI)
  Programmatically modify Fusion-series method files
Applications of IAPI
  Devin Schweppe – Fusion API
  Christoph Wichmann – Q Exactive API
Examples and FAQ for IAPI and XMMI
Applications of XMMI
  Pavel Shliaha – Fusion Method Creation
Notes: Overview of the first half of today
Some special guest presenters on their experience using these tools

## Slide 4: Thermo Fisher Scientific LSMS GitHub Page
  https://github.com/thermofisherlsms
Notes: All I will talk today will be on our github page

## Slide 5: Thermo Fisher Scientific LSMS GitHub Page
  https://github.com/thermofisherlsms
Mike Senko
Tuesday Poster
Covered Today
Covered Today
Notes: Please see Mike’s poster on Tuesday about programmable-xcalibur
I’ll cover these two repositories in today’s talk

## Slide 6: Overview
Instrument Application Programming Interface (IAPI)
  Real-time control and acquisition of Orbitrap Mass Spectrometers using .NET
XML Method Modification Interface (XMMI)
  Programmatically modify Fusion-series method files
Applications of IAPI
  Devin Schweppe – Fusion API
  Christoph Wichmann – Q Exactive API
Examples and FAQ for IAPI and XMMI
Applications of XMMI
  Pavel Shliaha – Fusion Method Creation
Notes: First I am going to give a brief overview of the IAPI and its capabilities

## Slide 7: Why an API?
“I wish I could do …”
“I have this awesome new experiment …”
“Can I do this in method editor?”
“I only want to select peaks that have odd masses …”
Notes: Why an API?
Lots of questions about doing cool things without mass spectrometers

## Slide 8: Why an API?
“I have this awesome new experiment …”
“I only want to select peaks that have odd masses …”
Users can only control the Mass Spectrometer through our Tune application and Method acquisition.
Although both provide great flexibility in acquisition strategies they don’t encompass all possibilities.
Providing an Application Programming Interface enables users to extend the capabilities of the Orbitrap mass spectrometers.
Allows real-time data acquisition and control.
“I wish I could do …”
“Can I do this in method editor?”
Notes: We provide both Tune and Method Editor to allow flexibility, but not everything is possible
By providing an API, we can extend the capabilities of the MS
It also allows real-time control of the MS

## Slide 9
The IAPI empowers users to programmatically control acquisition
of both the Exactive and Tribrid-series Mass Spectrometers
Notes: Here is just a quick demo of using a simple IAPI-enabled app to control the MS

## Slide 10: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started
Notes: Going to cover the following in the first talk
Will go over more examples, code, and questions in part 2

## Slide 11: Instrument API Git Repository
  https://github.com/thermofisherlsms/iapi
Public interfaces for
  Tribrid
  Exactive
Documentation
  Previous ASMS Posters
  Help files
VS Solution of Example Applications
External Requirements
Tune Installation
.NET 4.6.2+
License Agreement in place
IAPI License Key Activated
Notes: Again, the IAPI is hosted on our GitHub page
It contains all the interfaces and documentation to get started
You’ll need a few things to use it

## Slide 12: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started
Notes: I’ll cover the overall architecture of the API

## Slide 13: Interfaces
Exactive Series
Tribrid Series
IAPI Version
Instrument Extensions
Misc.
All interfaces are in the git repository, in the lib/ directory
Notes: We have two IAPI-enabled instruments lines
The Exactive series came first with the first API definitions
  It is broken into multiple assemblies.
The Tribrid series came next, with a newer API definition
  Here again, we have the extensions for the Fusion API

## Slide 14: Differences and Similarities between Exactive and Tribrid Series IAPI
Although IAPI 1.X is not binary compatible with IAPI 2.X, they share considerable amount of structure
  Namespaces and hierarchies are virtually unchanged
  A few name changes or minor reorganization
  Removal of some interfaces
  Spectrum-related interfaces moved into own assembly
Notes: I would like to highlight some of the differences between the exactive and tribrid APIs

## Slide 15: Differences and Similarities between Exactive and Tribrid Series IAPI
Although IAPI 1.X is not binary compatible with IAPI 2.X, they share considerable amount of structure
  Namespaces and hierarchies are virtually unchanged
  A few name changes or minor reorganization
  Removal of some interfaces
  Spectrum-related interfaces moved into own assembly
The logical flow for both receiving scan data and sending scans definitions to the instrument is identical
Notes: I would like to highlight some of the differences between the exactive and tribrid APIs

## Slide 16: Differences and Similarities between Exactive and Tribrid Series IAPI
Although IAPI 1.X is not binary compatible with IAPI 2.X, they share considerable amount of structure
  Namespaces and hierarchies are virtually unchanged
  A few name changes or minor reorganization
  Removal of some interfaces
  Spectrum-related interfaces moved into own assembly
The logical flow for both receiving scan data and sending scans definitions to the instrument is identical
Different methods for instantiating the IAPIs
  Tribrid-series uses a Factory method model with Microsoft’s Managed Extensibility Framework (MEF)
  Exactive-series uses the System.Reflection.Assembly API
Notes: I would like to highlight some of the differences between the exactive and tribrid APIs

## Slide 17: Differences and Similarities between Exactive and Tribrid Series IAPI
Although IAPI 1.X is not binary compatible with IAPI 2.X, they share considerable amount of structure
  Namespaces and hierarchies are virtually unchanged
  A few name changes or minor reorganization
  Removal of some interfaces
  Spectrum-related interfaces moved into own assembly
The logical flow for both receiving scan data and sending scans definitions to the instrument is identical
Different methods for instantiating the IAPIs
  Tribrid-series uses a Factory method model with Microsoft’s Managed Extensibility Framework (MEF)
  Exactive-series uses the System.Reflection.Assembly API
Some interfaces are not implemented due to technical reasons or deemed not immediately useful
Notes: I would like to highlight some of the differences between the exactive and tribrid APIs

## Slide 18
  The rest of this presentation will focus on the Tribrid-series IAPI.
  Most topics are applicable to the Exactive-series with minor differences.

## Slide 19: MS Architecture
Mass Spectrometer
Data System
Local Switch
Tune Application
Instrument
Service
Rawfile
Notes: Here is the basic lay out of the MS architecture

## Slide 20: IAPI Architecture
Mass Spectrometer
Data System
Local Switch
Tune Application
Instrument
Service
3rd Party
Application
IAPI-Enabled App
Rawfile

## Slide 21: IAPI Data and Control Flow
Handles scan Event
Processes the data
Do whatever you want!
Displays scan in Tune
Writes to Rawfile
Forwards data to API
Mass Spectrometer
Data System
Instrument
Service
3rd Party
Application
Scan Data
Stream
Converts data into IAPI format
Raises Event to any subscribers
Notes: If we start with the data flow of the spectrum through the service and API to the third party application

## Slide 22: IAPI Data and Control Flow
Calls IAPI to create IScanDefinition object
Changes definition as sees fit
Handles scan Event
Processes the data
Do whatever you want!
Displays scan in Tune
Writes to Rawfile
Forwards data to API
Mass Spectrometer
Data System
Instrument
Service
3rd Party
Application
Scan Data
Stream
Converts data into IAPI format
Raises Event to any subscribers
Validates scan definition
Forwards scan to Instrument Service
Scan Definition
Logs event
Downloads scan definition to instrument
Notes: We also allow the third party app to modify what the instrument is doing

## Slide 23: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started
Notes: Lets dig in deeper into the receiving side of things

## Slide 24: MS Scan Interfaces
IMsScanContainer exposes IMsScans
IMsScan contains metadata and derives from ISpectrum
ISpectrum contains a collection of ICentroids
ICentroids contains meta data and derives from IMassIntensity
IMassIntensity is the m/z and intensity data

## Slide 25: Scan Data Stream
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Receiving the scan data is through the Event Driven programming model
Scans are sent via the .NET Event and EventHandler mechanism
  Multiple handlers per scan can be registered
  Minimal overhead and latency
Receives Scans from both Tune and Method acquisitions

## Slide 26: Scan Data Stream
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Receiving the scan data is through the Event Driven programming model
Scans are sent via the .NET Event and EventHandler mechanism
  Multiple handlers per scan can be registered
  Minimal overhead and latency
Receives Scans from both Tune and Method acquisitions
tSIM running ~21 Hz (47.5 ms) from Tune.
  Average time of 48.4 ms between scans received (n = 200)

## Slide 27: Scan Data Stream
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Receiving the scan data is through the Event Driven programming model
Scans are sent via the .NET Event and EventHandler mechanism
  Multiple handlers per scan can be registered
  Minimal overhead and latency
Receives Scans from both Tune and Method acquisitions
tSIM running ~21 Hz (47.5 ms) from Tune.
  Average time of 48.4 ms between scans received (n = 200)
  The IAPI is responsive and has no impact on the MS acquisition

## Slide 28: Scan Data Contents

## Slide 29: Scan Data Contents

## Slide 30: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started

## Slide 31: IScans Interface
Sending scans to the instrument is through the IScans interface
  Provides methods for creating different types of IScanDefinitions
  Provides methods for setting and cancelling scans.

## Slide 32: IScanDefinition
Scans are defined in the IScanDefinition interface
  The RunningNumber property is for general purpose use
    It shows up in the Scan Trailer as the ‘Access ID’ field
  The Values property is a Dictionary<string,string>
    Keys are the scan parameters
    Values are their associated values
    Values can take different formats based on the parameter
    Invalid, or nonsense values, are ignored
Example of MS3 of 524@35 191@25
Scan 150 – 550 m/z

## Slide 33: IScanDefinition

## Slide 34: Scan Management
Mass Spectrometer
Data System
Local Switch
Instrument
Service
3rd Party
Application
IAPI-Enabled App
Sending Scan to MS

## Slide 35: Scan Management
Mass Spectrometer
Data System
Local Switch
Instrument
Service
3rd Party
Application
IAPI-Enabled App
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Sending Scan to MS

## Slide 36: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 37: Scan Management - ICustomScan
Single scan, drops to lower level after completion
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 38: Scan Management - ICustomScan
Single scan, drops to lower level after completion
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Once executed, waits the Single Processing Delay (s) for a new Custom scan to arrive*
Default delay is 0 seconds (no waiting)
Immediately ends the wait if a new Custom scan arrives or there are pending ones
  The Tribrid uses a circular buffer (40-capacity) to store them
* Single Processing Delay is not enabled when occurring during Tribrid method acquisition, due to acquisition pipeline

## Slide 39: Scan Management - IRepeatingScan
Repeats scan until explicitly canceled by user
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Continually runs the same scan until:
  Cancelled by the user
  Replaced by a new Repeating Scan
  Preempted by a Custom Scan

## Slide 40: Scan Management - IRepeatingScan
Runs Scans and logic defined in the method
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Runs the normal method logic until completion or new IAPI scan arrives

## Slide 41: Scan Management - IRepeatingScan
Runs the Scan defined in Tune
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 42: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method

## Slide 43: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating

## Slide 44: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom

## Slide 45: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating

## Slide 46: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating

## Slide 47: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating
Method Completed

## Slide 48: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Time
Start Method
Set Repeating
Set Custom
Canceled Repeating
Method Completed

## Slide 49: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority

## Slide 50: Scan Management
Scan Hierarchy
Tune
Method
Repeating
Custom
Priority
Method Scans
IAPI Scans
Method Scans

## Slide 51: Starting Acquisitions from the IAPI
The IAcquisition interface enables the IAPI to start/stop acquisitions
It also has Events for listening to when an acquisition starts and stops

## Slide 52: Starting Acquisitions from the IAPI
The IAcquisition interface enables the IAPI to start/stop acquisitions
It also has Events for listening to when an acquisition starts and stops

## Slide 53: IAPI Outline
IAPI Architecture
  Interfaces
  Data and Control Flow
Receiving Data from the MS
  Interfaces
  Scan Data Stream Subscription
Controlling the MS
  Sending Scan Definitions
  Changing other MS parameters
Getting Started

## Slide 54: License Activation
Access to the IAPI is given after a fully executed IAPI License is in force.
Contact ThermoMSLicensing@thermo.com to get started
  Please include which IAPI (Tribrid or Exactive) you would like access to in your initial correspondence

## Slide 55: License Activation
Access to the IAPI is given after a fully executed IAPI License is in force.
Contact ThermoMSLicensing@thermo.com to get started
  Please include which IAPI (Tribrid or Exactive) you would like access to in your initial correspondence
A license key will be provided to you and is entered in Tune’s “About Tune” dialog
Restart both the Instrument and Data system computer and the API will be enabled
Activating the License on the Tribrids

## Slide 56: Overview
Instrument Application Programming Interface (IAPI)
  Real-time control and acquisition of Orbitrap Mass Spectrometers using .NET
XML Method Modification Interface (XMMI)
  Programmatically modify Fusion-series method files
Applications of IAPI
  Devin Schweppe – Fusion API
  Christoph Wichmann – Q Exactive API
Examples and FAQ for IAPI and XMMI
Applications of XMMI
  Pavel Shliaha – Fusion Method Creation

## Slide 57
Method courtesy of Christopher Rose
The Method Editor offers tremendous flexibility in method creation;
sometimes manually editing complex methods is laborious.

## Slide 58: Programmatically Modify Fusion Methods
Method File
Modifications
Modified
Method File
XML Method
Modification Interface

## Slide 59: Method Modifications Git Repository
  https://github.com/thermofisherlsms/meth-modifications
Supports Fusion and Triple-Quad Methods
  Versions 1.2, 2.0, 2.1, 3.0 and 3.1
Documentation
Complete Examples
VS Solution
  Command Line Program
  Heavily commented code
Schemas
External Requirements
Tune Installation
  Full or Workstation
.NET 4.6.2+

## Slide 60: XML Modification Interfaces
Notes: Two .NET assemblies are needed:
Thermo.TNG.MethodXMLFactory.dll
Thermo.TNG.MethodXMLInterface.dll

## Slide 61: Complete Code Example

## Slide 62: Method Modifications Visual Studio Solution
The XmlMethodChanger Solution contains two projects
  XmlMethodChanger.lib
    A public, static class library (MethodChanger)
    Wraps the two Thermo.TNG.*.dll assemblies
    Provides public methods for common tasks
    Serves as an example of the API Usage
  XmlMethodChanger.cmd
    Command-line program that uses XmlMethodChanger.lib
    Uses third-party NuGet package (Command Line Parser Library)
    Can be used as a standalone tool for modifying methods

## Slide 63: Demonstration
The “template” method contains the structure of the experiment
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository

## Slide 64: Demonstration
The “template” method contains the structure of the experiment
The modification XML specifies a list of transformations to perform on the “template” method file
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository

## Slide 65: Demonstration
The “template” method contains the structure of the experiment
The modification XML specifies a list of transformations to perform on the “template” method file
The transformed method can then be saved to a separate .meth file
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository

## Slide 66: Demonstration
This example is from the examples/Fusion/IntensityThreshold/ directory in the Git repository

## Slide 67: Fusion 3.1 Method Modification Schema
3.1/MethodModification.xsd

## Slide 68: Method Modifications
The XML contains a collection of Modifications
Each Modification contains:
  An Order attribute
    This is the order that this modification will take effect, from low to high.
  Collection of Actions elements
    This is what the Modification performs on the method
    Experiments are performed first, in the order they are defined in the xml
    CopyAndAppendExperiment are performed second, in the order they are defined in the xml

## Slide 69: Copy And Append Experiment
Copies a method experiment and appends it the end of all the experiments
The attribute SourceExperimentIndex (0-based) specifies which experiment to copy

## Slide 70: Copy And Append Experiment
Copies a method experiment and appends it the end of all the experiments
The attribute SourceExperimentIndex (0-based) specifies which experiment to copy

## Slide 71: Experiment Modifications
Modifies an experiment based on the attribute ExperimentIndex (0-based)
Here you can target certain nodes/filters of the selected experiment and modify their contents
You can also modify the experiment time frame with the StartTimeMin and EndTimeMin fields.

## Slide 72: Source Node Position
SourceNodePosition is a mechanism to specify a specific node in scan tree

## Slide 73: Source Node Position
SourceNodePosition is a mechanism to specify a specific node in scan tree
Each instance of SourceNodePosition selects which child node to select; numbered left-to-right and 0-based.
0
1
0
1

## Slide 74: Source Node Position
SourceNodePosition is a mechanism to specify a specific node in scan tree
Each instance of SourceNodePosition selects which child node to select; numbered left-to-right and 0-based.
To specify the highlighted node, you would need to do:
0
1
0
1

## Slide 75: Source Node Position
SourceNodePosition is a mechanism to specify a specific node in scan tree
Each instance of SourceNodePosition selects which child node to select; numbered left-to-right and 0-based.
To specify the highlighted node, you would need to do:
  The first instance selects the left branch off the MS1.
0
1
0
1

## Slide 76: Source Node Position
SourceNodePosition is a mechanism to specify a specific node in scan tree
Each instance of SourceNodePosition selects which child node to select; numbered left-to-right and 0-based.
To specify the highlighted node, you would need to do:
  The first instance selects the left branch off the MS1.
  The second instance selects the right branch off the ddMS2
0
1
0
1

## Slide 77: Copy And Append Scan Node
Copies a scan node in the experiment and appends it to its parent node
Uses SourceNodePosition to specify the specific node in the tree

## Slide 78: Copy And Append Scan Node
Copies a scan node in the experiment and appends it to its parent node
Uses SourceNodePosition to specify the specific node in the tree

## Slide 79: Scan Node Modification
Modifies a scan based on the specified SourceNodePosition.
You can currently change the
  Isolation offset
  CID Collision Energy
  HCD Collision Energy

## Slide 80: Mass List Filter Modification
Modifies a filter with a mass list based on the specified SourceNodePosition and Above attribute
  To specify the filter, you first select the node the filter is attached to, and then indicate if the filter is above or below the node.
This filter would be defined as

## Slide 81: Mass Filter Modification
Modifies a filter with a mass list based on the specified SourceNodePosition and Above attribute
  To specify the filter, you first select the node the filter is attached to, and then indicate if the filter is above or below the node.
This filter would be defined as
And this filter would be defined as

## Slide 82: Mass List Modification
MassList contains a collection of MassListRecords which are the entries in the table
The different Boolean attributes control which columns in the table are used
The Append attribute indicates if the entries are append to the list or are overwritten

## Slide 83: Overview
Instrument Application Programming Interface (IAPI)
  Real-time control and acquisition of Orbitrap Mass Spectrometers using .NET
XML Method Modification Interface (XMMI)
  Programmatically modify Fusion-series method files
Applications of IAPI
  Devin Schweppe – Fusion API
  Christoph Wichmann – Q Exactive API
Examples and FAQ for IAPI and XMMI
Applications of XMMI
  Pavel Shliaha – Fusion Method Creation

## Slide 84: Overview
Instrument Application Programming Interface (IAPI)
  Real-time control and acquisition of Orbitrap Mass Spectrometers using .NET
XML Method Modification Interface (XMMI)
  Programmatically modify Fusion-series method files
Applications of IAPI
  Devin Schweppe – Fusion API
  Christoph Wichmann – Q Exactive API
Examples and FAQ for IAPI and XMMI
Applications of XMMI
  Pavel Shliaha – Fusion Method Creation

## Slide 85: IAPI FAQ
Is the IAPI Free?
  Yes! Though it requires a license agreement with Thermo.
Are the Q Exactive and Fusion IAPIs the same?
  They are not binary compatible, but share the same structure, namespacing, and basic data and control paths.
  It only requires a minimal amount of work to convert code between the two APIs.
How do we obtain the IAPI?
  Interfaces are shared on https://github.com/thermofisherlsms/iapi
  Contact ThermoMSLicensing@thermo.com to start the process.
Can you run the IAPI while acquiring an acquisition from a method?
  Yes, the IAPI works even when in idle, or actively running a method.

## Slide 86: IAPI FAQ
Are IAPI scans parallelized?
  When running a method, IAPI scans are inserted into the acquisition pipeline that parallelizes the scans with other scans.
  When running without a method, IAPI scans are not parallelized.
Are the results of the method filters published through the IAPI?
  The IAPI only publishes the raw spectrum, without any of the method filters taking effect.
Can we obtain profile data through the IAPI?
  Currently only the centroided data is returned to the IAPI.
Is the IAPI 64-bit enabled?
  No, the main instrument service is a 32-bit service and cannot export a 64-bit interface

## Slide 87: XMMI FAQ
Can you modify every parameter?
  No, only a subset of parameters in a method can be modified using this API.
I want to modify parameter X, Y and Z, can you provide it?
  Please file an issue with the github account describing the desired changes.
  https://github.com/thermofisherlsms/meth-modifications/issues
Do you need be attached to an instrument to modify methods?
  No, if you have installed the Workstation configuration, you can create and modify methods on any computer.
Does this work for Exactive methods?
  This interface does not currently support Exactive method modification.

## Slide 88: IAPI Example – Tribrid Series
This example is from the examples/tribrid/MinifiedExample/ directory in the Git repository
Use Thermo.TNG.Factory.Factory.Create() to create instrument access container
Connect to the main service using the StartOnlineAccess() call
Wait until server has connected and get access to a particular instrument

## Slide 89: Overview
Instrument Application Programming Interface (IAPI)
  Real-time control and acquisition of Orbitrap Mass Spectrometers using .NET
XML Method Modification Interface (XMMI)
  Programmatically modify Fusion-series method files
Applications of IAPI
  Devin Schweppe – Fusion API
  Christoph Wichmann – Q Exactive API
Examples and FAQ for IAPI and XMMI
Applications of XMMI
  Pavel Shliaha – Fusion Method Creation

## Slide 90: Controlling the Syringe from the IAPI
The ISyringePumpControl interface enables direct control of the syringe